Stepwise Guide for Biohazard Spills Inside and Outside the Biological Safety Cabinet in the BSL-3 Laboratory

Equipment:

TB Spill Kit

Contents:

Instructions (SOP) for clean up

20% Amphyl (1 liter bottle containing 200 mL of undiluted

Amphyl)  or appropriate disinfectant

Hair net

Shoe covers

Gown

N-95 respirators (1 sm/med, 1 med/lg)

Nitrile Gloves (2 pairs Med, 2 pairs Lg)

2 biohazard spill signs to place on the doors to the laboratory

Autoclave Bags- small and large

Absorbent paper towels and/or cotton

Forceps

When a spill occurs the laboratory staff must first decide how to react

The actions required are dependent on:

    - Where the spill occurred

      - Inside the BSC

      - Outside the BSC

    - The severity of the spill

      - Type of culture (solid or liquid)

      - Amount of live organisms involved

The least dangerous spills will usually occur inside the biological safety cabinet.  These spills are already contained in a biological safety cabinet and if the biological safety cabinet is working properly and the technologist is using the biological safety cabinet correctly they pose little risk to laboratory staff.

Once a spill occurs outside the biological safety cabinet the action that should be taken will depend on the type of spill. A minor spill is defined as the breakage of a single tube of solid medium containing Mycobacterium species, or dropping a plastic Petri plate growing Mycobacterium species so that the lid falls off (breakage may or may not occur), or as the spillage of the contents of a clinical specimen.

A major spill cannot be categorized exclusively by the volume or quantity of breakage or spill. More important is whether the material involved is a liquid, and, if so, what is the concentration of the Mycobacterium species in the suspension.

Steps for remediating a minor spill in the biological safety cabinet

|  | Actions | Comments |
| --- | --- | --- |
| 1. | Place absorbent paper towel or cotton over the spill and saturate with disinfectant such as 5% phenol/ hypochlorite | The spill kit should be located near biological safety cabinet, stocked with all the supplies needed for cleanup. A minor spill is the breakage of a single tube/plate of solid media |
| 2. | Allow to sit undisturbed for 20 minutes | Ample time must be allowed for the disinfectant to act |
| 3. | Place contaminated absorbent paper towel into a disposable bag while in the BSC, tie the bag and place it in a container to be autoclaved |  |
| 4. | Change gloves if gloves have been contaminated |  |
| 5. | Immediately wipe the interior surface and walls of the biological safety cabinet and any items or equipment within the cabinet with a paper towel/cloth saturated with a disinfectant solution (5% phenol). Follow with 70% alcohol to remove residue that could lead to corrosion of the biological safety cabinet. |  |
| 6. | File an incident report | Notify supervisor |

Steps for remediating a large spill inside the biological safety cabinet

|  | Actions | Comments |
| --- | --- | --- |
| 1. | Ensure the drain valve is closed | The drain valve should always be closed |
| 2. | Surface-decontaminate and remove all items from BSC | Ample time must be allowed for the disinfectant to act |
| 3. | Pour appropriate disinfectants onto the work surface |  |
| 4. | ALLOW AT LEAST 30 MINUTES FOR DECONTAMINATION! |  |
| 5. | Wipe down the tray (work surface), turn it over and wipe down the under side. Remove from cabinet after the appropriate disinfectant contact time. |  |
| 6. | Wipe down the front and back grates with ample disinfectant and remove them from the cabinet after appropriate contact time. |  |
| 7. | Pour disinfectant into the bottom pan of cabinet. ALLOW AT LEAST 30 MINUTES FOR DECONTAMINATION! |  |
| 8. | Empty the drain pan into a collection vessel containing disinfectants and autoclave. |  |
| 9. | Wipe the interior surface and walls of the biological safety cabinet and any items or equipment within the cabinet with a paper towel saturated with a disinfectant solution (5% phenol). Follow with 70% alcohol to remove residue that could lead to corrosion of the biological safety cabinet. |  |
| 10. | Replace and autoclave gown and respirator if either has become contaminated while cleaning up spill |  |
| 11. | File an incident report | Notify supervisor |

Steps to Remediate a minor spill outside the biological safety cabinet

|  | Actions | Comments |
| --- | --- | --- |
| 1. | Immediately cover the area with paper towels and completely soak with 20% amphyl solution or other appropriate disinfectant | Applying amphyl concentrically beginning at the outer margin and working toward the center. |
| 2. | Leave the room and place signs on the doors indicating a spill has occurred and the area should not be entered |  |
| 3. | Inform the Laboratory manager that a spill has occurred |  |
| 4. | Do not enter the room for at least 2 hours | This allows aerosols to settle and sufficient air exchanges to occur. |
| 5. | Put on the personal protective equipment found in the spill kit and double glove | Ensure sleeves of gown are tucked into gloves and respirator has made a good seal |
| 6. | Re-enter the laboratory and wipe up the spill placing all contaminated material in biohazard bags |  |
| 7. | Use forceps to pick up any broken glass and place it in a metal pan lined with a biohazard bag. |  |
| 8. | Label the pan “Broken Glass” | This enables the broken glass to be disposed of properly after autoclaving |
| 9. | Remove PPE and place in biohazard bag to be autoclaved | PPE should be removed in the following order: Outer gloves, respirator, gown, inner gloves. |
| 10. | Place all biohazard bags in the autoclave and start autoclave. |  |
| 11. | Wash hands with soap and water |  |
| 12. | Remove Biohazard spill signs from the doors | inform staff the room is safe to reenter |
| 13. | Complete and file incident report form | Notify supervisor |

Steps to remediate a major spill outside the Biological Safety Cabinet

|  | Actions | Comments |
| --- | --- | --- |
| 1. | IMMEDIATELY evacuate the room. Ensure biological safety cabinets and room exhaust fans are running. Staff should remove PPE while quickly walking toward the door, dropping it on the floor just inside the door when exiting | The aerosols produced from such a spill are more dangerous than any need to cover the spill. All affected areas should be evacuated. |
| 2. | Place a Biohazard sign on the door that lead to the TB laboratory. This sign is found in the biohazard spill kit | This sign must indicate that a spill has occurred and that area should not be entered |
| 3. | Notify the laboratory manager or supervisor immediately |  |
| 4. | Allow the room to ventilate for 4 hours | this will dependent on the number of airchanges/hour in the BSL3. If the spill involved large volumes of liquid allow the room to ventilate overnight |
| 5. | Wearing protective clothing (protective gown/overalls, shoe covers, hair cap, gloves, and respirator), clean up the spill |  |
| 6. | Cover the spill with paper towels to soak up large volumes of liquid |  |
| 7. | Fill the bottle containing the undiluted amphyl (from the spill kit) with water to the 1L mark. Gently pour amphyl over the absorbent materials and the spill area | Amphyl is stored in its undiluted form because it will remain active longer then when diluted |
| 8. | Allow to stand undisturbed for 20 minutes |  |
| 9. | Place the absorbent material in a biohazard bag (fasten securely) or a lined discard pan | CAUTION: Use forceps or tweezers to pick up any broken glass and place it in a separate discard pan labeled “Broken Glass” so that it can be properly discarded after autoclaving. |
| 10. | Remove the protective clothing in the correct order and place in autoclave bags and autoclave with the discard pans/biohazard bags that contain the contaminated paper towels used to clean up the spill. |  |
| 11. | Remove the Biohazard Spill sign from the door once the spill is cleaned up and the room is safe to enter |  |
| 12. | Complete and file incident report form | Notify supervisor |

Steps to flush an eye splash

| 1. | Immediately flush eye at the eyewash station with water for an extended period of time (at least 5 minutes) |  |
| --- | --- | --- |
| 2. | Seek medical attention |  |
| 3. | Notify the supervisor and manager (or their designee) of the incident |  |
| 4. | Fill out an incident report form |  |