TABLE OF CONTENTS

1.	PURPOSE	2

2.	SCOPE	2

3.	RESPONSIBILITIES	2

4.	CROSS-REFERENCES	2

5.	PROCEDURES	3

5.1.	General Safety	3

5.2.	Set up	3

5.3.	Operation	3

5.4.	Calibration	3

5.5.	Maintenance	4

5.5.1  Cleaning	4

5.5.2. Servicing	5

6.	REFERENCES	5

7.	CHANGE HISTORY	5

# PURPOSE

This SOP describes the using and maintenance of the heating block.  The heating block is used for a range of procedures, including heat killing of mycobacteria during specimen processing for molecular methods.

# SCOPE

This SOP covers procedures requiring use of the heating block, located in the Specimen Processing laboratory, performed in the FIND Research Laboratory, Kampala, Uganda.

# RESPONSIBILITIES

All staff members working in the FIND Uganda Research Laboratory are responsible for the implementation of this operating procedure.

All users of this procedure who do not understand it or are unable to carry it out as described are responsible for seeking advice from their supervisor.

# CROSS-REFERENCES

Document Matrix_Uganda QP 01-03-03

Refer to SOPs listed under 07-01 (General Procedures) and 07-05 (Molecular Methods).

| See: | Document Matrix_Uganda QP01-03-03.doc |
| --- | --- |
| Location: | Hard copy: FIND Uganda SOPs |

# PROCEDURES

## General Safety

- Do not use heating block in an explosive environment, or in the presence of a flammable or combustible material.

- Do not heat substances that react violently to heating.

- Do not touch block when unit is hot or heating.

- Avoid spilling liquids into wells or into unit ventilation holes.

- Connect to properly grounded outlet.

## Set up

- Install heating block after checking compatibility of voltage supply.

- Ensure that heat block is placed on a firm and flat surface.

- Select location that is dry and free of moving air from heating or air conditioning vents.

- Ensure that unit is placed at minimum of six inches away from all other surfaces.

## Operation

- Turn unit on with On/Off switch at rear of unit.

- Use Up and Down arrow keys to set desired operating temperature.

- Press Start to begin heating unit to set point.

- On initial heating, the display will show a temperature overshoot that results from the rapid heating of the block, but the actual block and sample temperature do not overshoot.

- To change temperature setpoint, press Stop to reenter temperature set mode before attempting to change set point with Up and Down arrows.

## Calibration

- Digital Dry Baths are calibrated at the factory at 37°C using a small hole pattern block. If using a platform style block or a block with large holes, one may consider recalibrating unit.  One may also consider recalibrating unit if using loosely fitting or strangely shaped reaction vessels.

- To calibrate, turn unit off with On/Off switch.

- Place thermometer of known accuracy into block thermometer hole.

- Press and hold Start/Stop key while simultaneously using On/Off switch.

- Block will emit sound and display will flash a set temperature.

- Press Up and Down arrows to set desired temperature. Then press Start.

- Allow 40-50 min for unit to heat to set temperature and equilibrate there.  The entire LED display will flash upon equilibration.

- After equilibration, read thermometer and use Up and Down arrows to adjust display to thermometer reading.  Then press Start-Stop key.

- Unit will begin to adjust operating temperature to original setpoint while including recalibration factor.  Allow sufficient time for re-equilibration and then again compare temperature on thermometer to the displayed temperature.  Repeat procedure if required.

## Maintenance

### 5.5.1  Cleaning

- Ensure that unit is dry and cool and that the power supply is disconnected before performing any cleaning or maintenance.

- If any liquid spills into well or bath, disconnect plug from outlet and turn unit upside down to minimize liquid contacting internal components.  Remove bottom cover and inspect to assure that liquid has not contacted heater elements, electronic controls, or connectors. Use only qualified service technician to clean unit and replace damaged parts.

- Unit may be cleaned with moist cloth containing mild soap solution.  Unit should never be immersed in any liquid.

- Heating blocks may also be cleaned with mild soap solution.

- Record cleaning in the Laboratory Cleaning and Maintenance Logbook.

| Use: | Laboratory Cleaning and Maintenance Logbook_form.doc |
| --- | --- |
| Location: | Hard copy: FIND Uganda SOPs |

### 5.5.2. Servicing

- In case of damage or failure to work, an Equipment Failure Report must be completed and the authorized service representatives must be called out.

- An Equipment Failure Notice must be placed on the thermal cycler indicating that work is prohibited.

| Use: | Equipment Failure Notice _sign.doc |
| --- | --- |
| Location: | Hard copy: FIND Uganda SOPs |

| Use: | Equipment Failure Report _form.doc |
| --- | --- |
| Location: | Hard copy: FIND Uganda SOPs |

# REFERENCES

Reference Manual: Digital Dry Bath Models D1100 and D1200 by Labnet

# CHANGE HISTORY

| New version # / date | Old version # / date | No. of changes | Description of changes | Source of change request |
| --- | --- | --- | --- | --- |